Motor Trouble Shooting

(Note: The term 115 volts is used as a generic voltage reading. It is acceptable to have 110 to 125 volts. The following trouble shooting will request electrical tests from wire to wire. These are not tests from wires to ground.)

- If the motor hums but does not turn check the following items:

- Check the voltage between the red and white wires on the mercury relay with the machine power on but not running a cycle. Then do the same test while starting the machine. We are looking for a drop in voltage. If the drop in voltage between the two tests is close to 10 % you will need to correct this problem first. It could be an incoming power problem. A drop in voltage can keep a motor from starting and will damage it.

If you have 115volts at the black and white wires on the mercury relay while the machine is running a cycle and did not find a voltage drop then check to see if the motor shaft will turn freely. This can be checked where the shaft comes out of the motor in front or by removing the round cover on the center of the back of the motor. Check with machine off.

If the shaft will not turn, open the pump at the gasket and look for the cause.

If the shaft turns freely start the cycle and check for 115 volts at the black and white wires in the back of the motor that come from the control box.

If you have 115 volts at the motor wires and the spade connectors are tight and all the wires are on the correct terminals then change the motor.

- If the motor does not hum or runs check the following items:

  - Start the cycle and check for 115 volts on the mercury relay. You should have 115 volts between the brown and the white wires, the red and the white wires, and the black and the white wires.

If you do not have 115 volts between the brown and the white wires, follow the brown wire until you find the problem.

If you do not have 115 volts between the red and white wires, follow the red wire until you find the problem.

If you do not have 115 volts between the black and white wires but you do have 115 volts between the red and white wires and the brown and white wires then change the mercury relay.

If you have 115 volts on all three of the wire combinations mentioned above then you should check for 115 volts at the black and white wires connected to the motor while the machine is running a cycle.

If you have 115 volts at the black and white wires at the motor, the spade connectors are tight and all the wires are hooked to the motor in the proper positions and it still won’t run then change the motor.

VOLTAGE IS A MOTORS BEST FRIEND OR WORST ENEMY

Low voltage will cause a motor to not start, blow a breaker, or fail prematurely. The low voltage causes the motor to stay connected to the start winding too long or to rapidly switch between start and run windings. This will reduce the life of the start windings and put excessive wear on the start switch.

Low voltage is a term used to describe a voltage reading that is below standard voltage or a voltage reading that drops too much when put under a load.

Low voltage readings taken while a machine is not running are usually caused by too many items on a circuit. This is why dishwashers should always be on their own breaker. This problem may also occur because too many items are connected to the same breaker panel, even if they are on separate breakers.

Voltage drops or brown outs can be caused by the same items listed above or too small of a circuit. This is one of the few times that electricity and water are alike. If you have 100 psi water pressure at the machine but you are using a ¼ inch water line, the machine will not fill fast enough. If you have 115 volts at the machine but the wires are too small or other items are using the same circuit, the flow of the electricity will be reduced which causes the low voltage-brown out situation.

REMEMBER

1. Do not over look the easy fix.

2. Do not always assume it is a motor.

3. Do not demote yourself to a parts changer. We are dishwasher technicians.

4. A loose connection can make a machine act like it has a bad motor but is a lot easier to fix than changing a motor.

5. The faster we get the customer up and running the happier they are and the better we look.